XMM Optical Monitor

MULLARD SPACE SCIENCE LABORATORY

UNIVERSITY COLLEGE LONDON	Author: Nigel R Bray

MSSL  XMMOM

DOCUMENT:

Distribution:

| XMM-OM Project Office | R Hunt | ü |
| --- | --- | --- |
| ESA PX | H Eggel |  |
|  | S Thürey |  |
| CSL | P Rochus |  |
|  | S Roose |  |
| Pennsylvania State University | S Horner |  |
| Royal Greenwich Observatory | R Bingham |  |
| University College London | J Fordham |  |
| Mullard Space Science Laboratory | A Dibbens | ü |
|  | D Self |  |
|  | G Davison |  |
|  | H Huckle |  |
|  | H Kawakami | ü |
|  | M Carter | ü |
|  | M de la Nougerede |  |
|  | M Whillock |  |
|  | N Bray |  |
|  | P Smith |  |
|  | P Thomas | ü |
|  | R Card |  |
|  | R Chaudery |  |
|  | T Kennedy |  |
|  | T Patrick |  |
|  | K Mason |  |
|  | A Smith | ü |

| Author: | NR Bray | Date: | 18-Nov-96 |
| --- | --- | --- | --- |
| OM Project Office |  | Date: |  |
| Distributed: |  | Date: |  |

| Issue | Date | Changes |
| --- | --- | --- |

| Machine File-System Directory: | \\mssli7\data\projects\ |
| --- | --- |
| Document File Name: | input.doc |
| Application: | MS Word 7 |
| Included files: |  |
| Print Date (this): | d-MMM-yy 35387.7694444444 HH:mm |
| Printer: | Postscript |

# 1.	Scope

This document contains the requirements for the supply of 5 science and space flight grade imaging Charge Coupled Devices within 6 months from the placement of an order.

# 2.	Reference Documents

[1]	ASTM F 672-88, Spreading resistance in epitaxial layers.

[2]	ESA/SCC 20490, Internal Visual Inspection of Integrated Circuits

[3]	ESA/SCC 20500 External Visual Inspection

[4]	ESA/SCC 21400. Scanning Electron Microscope Inspection Procedures

[5]	MIL-STD-883, Test Method 1008, Test Condition D

[6]	MIL-STD-883, Test Method 1010, Test Condition B

[7]	MIL-STD-883, Test Method 2001, Test Condition A

[8]	MIL-STD-883, Test Method 2002, Test Condition B

[9]	MIL-STD-883, Test Method 2007, Test Condition A

[10]	MIL-STD-883, Test Method 2011, Test Condition D

[11]	MIL-STD-883, Test Method 2023, Test Condition D

[12]	NAS 3447 Contamination Barrier for Precision Cleaned Items

# 3.	Technical Requirements

## 3.1	Functional

R1	Frame Transfer mode

R2	Image region  385 x 288 pixels

R3	Storage region  385 x 290 pixels

R4	Non-anti-bloomed

## 3.2	Mechanical

R5	Package: 20 pin 15.24mm pitch DIL

R6	Fibre stub attached to CCD surface strength 100G @ 4.5g

R7	The thickness of the package including fibre stub shall not exceed 4.2mm

R8	Position accuracies  of bonded stub
Parallelism: <3um edge to edge of CCD
Azimuth: +-1 degree
Overlap: min 1mm of stub outside CCD image region

## 3.3	Electrical

R9	Horizontal pixel shift frequency 10 MHz

R10	Vertical frame shift frequency 2 MHz

## 3.4	Thermal

R11	Nominal Temperature 50 °C at which optical requirements shall be met unless otherwise stated.

R12	Storage Temperature -50 to +85 °C

R13	Operating Temperature 0 to +65 °C

## 3.5	Optical

R14	22 +- 2um Square pixels

R15	Defects
Defect defined as +-2% of average pixel value with filter block
Integration time of 12ms @ Nominal Temperature
Column: none
White spots: none with size >2 pixels
Black Spots: <5
Horizontal: none

R16	Defect measurement conditions
Illumination equivalent to half saturation

R17	CTE > 99.998%

R18	QE > 30% @ 520nm

R19	Full well capacity shall be >300 000 electrons.

R20	Average dark signal <50 000 electrons per pixel per 120ms @ 50°C

R21	Standard deviation of dark current over all pixels <5% of average dark signal

R22	Conversion factor >0.8mV/electron

R23	There shall be no hot pixels, defined as >50 000 electrons above the average dark signal with an integration time of 120ms @50°C

R24	There shall be <5 warm pixels, defined as >10 000 electrons above the average dark signal with an integration time of 120ms @50°C

R25	There shall be no black pixels, defined as <0.95 times average signal.

R26	There shall be no hot columns, defined as >1000 electrons/pixel above average dark signal.

R27	There shall be no black columns, defined as <0.98 times average signal.

# 4.	Product Assurance Requirements

## 4.1	Manufacturing and Test Flow

MANUFACTURING AND TEST FLOW

µ §

R28	Wafer epitaxial layer spreading resistance shall be within 2kW.cm and 8kW.cm at 25um from the surface measured in accordance with [1] ASTM F 672-88, Spreading resistance in epitaxial layers..

R29	The thickness of the epitaxial layer shall be 50 +- 5um

R30	Prior to wafer manufacturing one wafer per epitaxy lot shall be measured for the spreading resistance and layer thickness of the epitaxial layer.  All data shall be recorded.  Wafers that do not meet the requirements above shall be rejected.

R31	After full wafer manufacturing and wafer sawing, a scanning electron microscope inspection shall be carried out on samples from each wafer lot in accordance with the inspection procedure [4] ESA/SCC 21400. Scanning Electron Microscope Inspection Procedures and keeping reference photographs available for later review.

R32	A sample of 3 test structure devices from each wafer manufacturing lot shall be tested with a Co60 source for radiation damage.

R33	Prior to the assembly of the chip to the support wafer chips shall be inspected for visual defects according to [2] ESA/SCC 20490, Internal Visual Inspection of Integrated Circuits and rejects removed from the lot.

R34	One sample each of 3 wafers from each epitaxy lot shall be measured for spreading resistances of the epitaxial layer. If the 3 samples do not give consistent results within 10%, further samples shall be taken such that wafers not meeting the spreading resistance requirements are rejected from the flight lot.

R35	From each assembly lot, 3 sample devices shall be tested for destructive wire bond pull in accordance with [10] MIL-STD-883, Test Method 2011, Test Condition D prior to any further production wire bonding.

## 4.2	Production Testing

PRODUCTION TESTS

µ §

R36	Testing and conditioning shall be carried out in an inert atmosphere (dry nitrogen or vacuum).

R37	The mechanical dimensions of packaged devices shall be measured and rejects removed from the lot.

R38	The position and alignment of the CCD with respect to the package shall be measured and recorded.

R39	All wire bonds shall be tested non-destructively in accordance with Error: Reference source not found.

R40	High temperature stabilisation bake shall be performed in accordance with [5] MIL-STD-883, Test Method 1008, Test Condition D during 48 hours at 125°C.

R41	Thermal cycling shall be applied in accordance with [6] MIL-STD-883, Test Method 1010, Test Condition B.

R42	Vibration testing shall be performed on 3 samples, after all assembly and conditioning, from each assembly lot in accordance with [9] MIL-STD-883, Test Method 2007, Test Condition A but at 30G.  The lot is passed only if there is no evidence of mechanical damage.

R43	Pre burn-in testing shall be defined and performed in order to select devices to meet the post-burn in tests and screening.

R44	Devices shall be visually inspected in accordance with [2] ESA/SCC 20490, Internal Visual Inspection of Integrated Circuits and [3] ESA/SCC 20500 External Visual Inspection.

## 4.3	Burn-In and Screening

BURN-IN AND SCREENING

µ §

R45	Burn-In and Screening shall be carried out on devices that have successfully passed the final production tests.

R46	All delivered devices shall undergo burn-in and screening.

R47	All conditioning, testing and storage of devices shall be in an inert atmosphere (dry nitrogen or vacuum)

R48	Burn-in conditions shall be:
Ambient temperature: 125°C
Duration:	168hrs

R49	Electrical and Optical test measurements shall be performed after burn-in (post burn-in tests) according to the details given in Section Error: Reference source not found “Error: Reference source not found”.

R50	Post burn-in tests shall be performed within 240 hrs.

R51	Devices shall be subject to a final visual inspection after the fibre block has been attached in accordance with [2] ESA/SCC 20490, Internal Visual Inspection of Integrated Circuits and [3] ESA/SCC 20500 External Visual Inspection.

## 4.4	Lot Acceptance Testing

R52	Lot acceptance testing shall be carried out on devices that have successfully passed the burn-in and screening and had the fibre block attached.

R53	The sequence and sample sizes for lot acceptance testing shall be as specified in the chart below:

LOT ACCEPTANCE TESTS

µ §

R54	All samples shall have gone through burn-in and screening.

R55	The shock tests shall be performed according to [8] MIL-STD-883, Test Method 2002, Test Condition B.

R56	The vibration tests shall be performed according to [9] MIL-STD-883, Test Method 2007, Test Condition A.

R57	The constant acceleration tests shall be performed according to [7].

R58	Shock, vibration and constant acceleration tested devices shall undergo a basic functional and performance check that assures that the devices still meet the post burn-in test limits.

R59	Accelerated life test conditions shall be:
Ambient temperature: 125°C
Duration:	1000hrs

R60	Accelerated life tested devices shall undergo the post burn-in tests and pass the performance criteria.

R61	The destructive wire bond pull tests shall be performed according to [10] MIL-STD-883, Test Method 2011, Test Condition D with post burn-in levels.

R62	All bonds in the device shall be tested.

## 4.5	Marking

R63	Each device shall be clearly marked with an identifier that is uniquely referenced to the device’s data pack.

## 4.6	Packaging and Transportation

R64	The external packing shall consist of:
An aluminium container with shock absorbing material that may contain several packaged devices.

R65	Devices shall be packed individually such as to avoid any damage and deterioration from the transport and storage environment.

R66	Individual device packing shall consist of:
An outer heat sealed nylon bag meeting the requirements in Error: Reference source not found Class 3.
An inner heat sealed polythene bag meeting the requirements in Error: Reference source not found Class 1, containing sufficient humidity absorber for 3 months transport and storage.

# 5.	Test Methods and Procedures

## 5.1	Electrical and Optical Test Methods

1. 	Integrate 100 frames digitally, full area mode

2. 	As 1. But vertical clocks stopped, horizontal clock stopped while integrating

3. 	As 1. But each line is summed after readout to give 1d result, 1 value per column

4. 	As 1. But clock an additional row and column

5. 	Derive photon transfer curve

## 5.2	Optical Test Measurements and Limits

|  |  |  |  |  |  | Limits |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test No | Parameter | Symbol | Test Method | Notes | Conditions | Min | Max |  | Units |  |
| 1 | Average dark signal | SD | 1 |  | 1 |  | 5000 |  | e- |  |
| 2 | Hot pixels in image and store | NHP | 1 |  | 1 |  | 0 |  |  |  |
| 3 | Warm pixels in image | NWI | 1 |  | 1 |  | 5 |  |  |  |
| 4 | Warm pixels in store | NWS | 1 |  | 1 |  | 5 |  |  |  |
| 5 | Hot pixels in output register | NHO | 2 |  | 1 |  | 0 |  |  |  |
| 6 | Black pixels in image | NBI | 1 |  | 2 |  | 0 |  |  |  |
| 7 | Black pixels in output register | NBO | 2 |  | 2 |  | 0 |  |  |  |
| 8 | Hot columns | NHC | 3 |  | 1 |  | 0 |  |  |  |
| 9 | Black columns | NBC | 3 |  | 2 |  | 0 |  |  |  |
| 10 | Charge transfer efficiency, image and store | CTEV | 4 |  | 4 |  | 99.998 |  | % |  |
| 11 | Charge transfer efficiency, output register | CTEH | 4 |  | 4 |  | 99.998 |  | % |  |
| 12 | Charge to voltage conversion factor | CVF | 5 |  | 5 |  | 0.8 |  | mV/ e- |  |
| 13 | Output reset noise | SN | 1 |  | 1 |  | 120 |  | e- RMS |  |
| 14 | Reset feed through voltage | VRESET | 1 |  | 1 |  | 80 |  | mV |  |
| 15 | Dummy output time constant | tD | 1 | 1 | 1 |  | See Note |  | ns |  |
| 16 | Signal output time constant | tO | 1 |  | 1 |  | 20 |  | ns |  |

Notes:

1. 	±10% of signal output time constant.

Conditions:

1. 	1MHz, 120ms, 50°C, Dark

2. 	1MHz, 120ms, 50°C, Illumination to give 100 000 electrons average signal

3. 	1MHz, 120ms, 10°C, Dark

4. 	1MHz, 120ms, 10°C, Illumination to give 100 000 electrons average signal on end pixels only

5. 	1MHz, 120ms, 10°C, Illumination various levels 100 to 200 000 electrons

## 5.3	Electrical Test Measurements and Limits

|  |  |  |  |  |  | Limits |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Test No | Parameter | Symbol | Test Method | Notes | Conditions | Min | Max |  | Units |  |
| 1 | Leakage current on input gates | IC | SCC 25000 § 5.1 |  | 15V |  | 5 |  | pA |  |
| 2 | Substrate through resistance | RSS | DC Test | 1 | 10mA |  | 100 |  | W |  |
| 3 | Forward diode voltages | VF | DC Test |  | 1mA | 0.2 | 1.5 |  | V |  |
| 4 | Reverse diode leakage currents | IR | DC Test |  | 24V | 50 |  |  | mA |  |
| 5 | Transistor drain to gate leakage currents | IDR | DC Test |  | 24V | 50 |  |  | mA |  |
| 6 | Gate to gate leakage currents | IGG | DC Test |  | 20V | 30 |  |  | mA |  |
| 7 | Gate protection voltages | VGP | DC Test |  | 30mA | 19 | 27 |  | V |  |
| 8 | Gate input capacitance | CÆ | DC Test |  | 10V |  | 2500 |  | pF |  |
| 9 | Gate to gate capacitance | CÆM | DC Test |  | 10V |  | 1000 |  | pF |  |
| 10 | Output register capacitance | CO | DC Test |  | 10V |  | 150 |  | pF |  |
| 11 | Reset input capacitance | CR | DC Test |  | 10V |  | 10 |  | pF |  |

Note 1 : Measured at wafer level only.

## 5.4	LAT Performance Test

## 5.5	LAT End of Life Test

# 6.	Programme Requirements

## 6.1	Deliverables

R67	Five parts from an assembly lot that has passed all testing and screening requirements shall be delivered.

## 6.2	Schedule

R68	Parts shall be delivered within 6 months of the placement of an order.

## 6.3	Documentation

R69	Documentation as defined here shall be made available at the delivery review prior to delivery and shipment.

R70	All documentation packages shall be identified with
Reference to device specification and issue
Reference to purchase order and date
Device serial numbers, assembly lot, and wafer manufacturing lot.

R71	A data pack shall be provided for each device including
Device wafer traceability
Production test measurements
Post burn-in test measurements

R72	A data pack shall be provided for each wafer manufacturing lot including
Wafer production report

R73	A data pack shall be provided for each assembly lot including
Device and wafer traceability
Failure analysis reports
Lot acceptance test report

Requires no pixel mask on end of row.

If present on the device.