XMM Optical Monitor

| MULLARD SPACE SCIENCE LABORATORY | Authors: T E Kennedy, H.E.Huckle, P.J.Smith, M.C.R.Whillock |
| --- | --- |
| UNIVERSITY COLLEGE LONDON |  |

XMM-OM (FM) INTEGRATED FUNCTIONAL TEST PROCEDURE

Document Number:  XMM-OM/MSSL/SP/0194.2

Distribution:

| XMM-OM Project Office | R Hunt | Orig. |
| --- | --- | --- |
| ESA PX | H Eggel |  |
| CSL | P Rochus |  |
|  | S. Roose |  |
| UCSB | T Sasseen |  |
| Royal Greenwich Observatory | R Bingham |  |
| University College London | J Fordham |  |
| Mullard Space Science Laboratory | R Card |  |
|  | M Carter |  |
|  | R Chaudery |  |
|  | A Dibbens |  |
|  | J Holmes |  |
|  | H Huckle |  |
|  | H Kawakami |  |
|  | T Kennedy |  |
|  | T Patrick |  |
|  | D Self |  |
|  | P Smith |  |
|  | P Thomas |  |
|  | M Whillock |  |
|  | J Lapington |  |
|  | K Mason |  |
|  | A Smith |  |

| Author: |  | Date: |  |
| --- | --- | --- | --- |
| OM Project Office |  | Date: |  |
| Distributed: |  | Date: |  |

CHANGE RECORD

| Issue | Date | Comments |
| --- | --- | --- |
| 1 | 29 July ‘98 | 1st Issue |
| 2 | 31 July ‘98 | Corrected aliases Clarified f/w actions Clarified HV actions Updated currents |

Contents

1.0 Scope

2.0 Documentation

3.0 General requirements

4.0 Test Procedures

5.0 Switch-on Test

5.1 Basic Mode Test

5.2 Operational Mode Test

5.3 Filter Wheel Control Test

5.4 Dichroic Control Test (beam deflector)

5.6 Image Acquisition Test - Multiple Window

5.7 Image Acquisition Test - Full Frame

5.8 Switch-off

Tables

Table 01: OM power consumption (with OM-1 secondaries on)

Table 02  Heater switching power profile

Table 03: Pref look-up Table

# 1.0	Scope

The test procedure is applicable to the flight model optical monitor test during the spacecraft integration and test phases.

The procedures will be conducted from the spacecraft EGSE, whereby the OM instrument station will follow the telecommand and telemetry transfer to and from the instrument. Analysis of science data will be performed on the OM instrument-station.

# 2.0	Documentation

- Applicable documents

- XMM-OM User manual Part 3 (XMM-OM/MSSL/ML/0010)

- XMM-OM User manual Part 4 (XMM-OM/MSSL/ML/0011)

# 3.0	General requirements

- Environmental conditions

  - No specific environmental conditions are envisaged other than those applicable for the XMM satellite.

- Safety precautions

  - The High Voltage Unit should not be switched on without consultation with the OM team.

  - All packets of type, subtype 3,2, 3,3, 4,2 and 4,3 should invalidate the test.

# 4.0	Test Procedures

The table below provides an overview of the test procedure to be used for the tests at spacecraft level.

The procedures are applicable to the flight model of XMM-OM only and should be applied to both prime and redundant.

| Description | To be performed | Test Sequence | Procedure |
| --- | --- | --- | --- |
| Interface test Estimated test time 20 minutes | AFT | Switch-on Boot mode test Switch-off | 5.0 (ehmod_pwr_on) 5.1 (ehmod_boot_test) 5.8 (ehmod_pwr_off) |
| Short Functional Test Estimated test time 30 minutes |  | Switch-on Boot mode test Switch to operational Filter wheel test Heater test Switch-off | 5.0 5.1 5.2 (ehist_operation) 5.3 (ehist_fwl_test) 5.5 (ehist_htr_test) 5.8 |
| Full Functional Test Estimated test time 240 minutes |  | Switch-on Boot mode test Switch to operational Filter wheel test Dichroic test** Heater test Image Test (multiple window) Image test (full frame) Switch-off | 5.0 5.1 5.2 5.3 5.4 (ehist_dic_test) 5.5 5.6 (ehist_img_multi) 5.7 (ehist_img_full) 5.8 |

** The Dichroic mechanism is a limited life system and must not be rotated more than 50 times.

The following section provides a detailed description of the test sequences to be used for testing of the optical monitor instrument.

Pass/Fail criteria

The following pass/fail criteria are applicable for all the tests performed on the OM:

- If one of the TC error packets is received the test shall be halted followed by a 	review board.

- TM parameters mentioned in the procedure will be checked. If not correct the test 	will be halted followed by a review board

- The test will be stopped if any TM parameter exceeds the defined limits.

- If an unexpected event TM packet is received the test shall be halted  followed by 	a review board.

# 5.0	Switch-on Test

WARNING

The Filter wheel is a limited life time device. Each rotation of the filter wheel (which occurs every time the OM is switched on) must be recorded in the instrument log book for the flight model.

Description

Switch-on procedure for the OM prime or redundant.

Switch-on the OM power and keep-alive line and check for HK data.

Initial state

The OM is switched off (KAL can be on).

Final State

The OM (prime or redundant) is switched on and in basic mode.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial State: OM 28V power switched off KAL line on |  |  |  |  |
| 1 1.1 1.2 1.3 | Switch-on OM power Switch-off the S/C heater (if not switched off before) Switch-on 28V Switch-on KAL (if not switched on before) |  | Filter wheel to coarse S/C procedure S/C procedure S/C procedure |  |  |
| 2 2.1 2.2 2.3 | Check Housekeeping Ensure HK monitor prog working Verify basic mode is operational Verify heaters are on Secondaries disabled | 90000 | (HK Valid after 15 second) pref H5395 = ‘Initial’ pref H5240 =’ononoffoff’ pref H5255 = ‘Disabled’ |  |  |
| 3 3.1 3.2 3.3 | OM power measure OM +28V voltage and current Measure OM KAL voltage and current Compare with table 1 |  | S/C procedure report: voltage and current S/C procedure report: voltage and current |  |  |
| 4 | Exit |  |  |  |  |

| Power line | Voltage | Current |
| --- | --- | --- |
| +28V main power | +28V +/- 2V | 1850mA +/- 20mA |
| Keep Alive Line | +8V +/- 1V | <18mA |

Table 01: OM power consumption (basic mode)

## 5.1	Basic Mode Test

Description

The function of the OM is tested in the basic mode test without loading the main application software (operational mode).

Dump IC memory and verify the operation of the KAL memory.

Initial state

The OM (prime or redundant) is switched on and in basic mode.

Final state

The OM (prime or redundant) is switched on and in basic mode.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial State: OM is switched on and in basic mode KAL power line is on |  |  |  |  |
| 1 1.1 1.2 1.3 1.4 | Check test command Send invalid command (tc_send) Check packet response Check HK (bad TC counter) Send valid command (tc_test) Check packet response Check HK | TC(0,0) 91202 90000 H1 91100 90000 | TM (3,2) Pref H5390 = 1 TM (3,1) Pref H5385 = 1 |  |  |
| 2 2.1 2.2 2.3 2.4 | Dump and verify KAL is OK Dump instruction space Address 0X3800 = Start Length = 16 (10 Hex) Words Compare contents: 85D0, 4964, 8320, 0906, B122 4820, 400E, 4800, 4009, 8520, FD40, 4820, 2000, 4810, 2002, B700, 0D00, 0D02, 0D04, 85F0 | H4101 94201 |  |  |  |
| 3 | IF memory contents is O.K the goto section 5.2, step 1.3. If memory contents is incorrect goto section 5.2, step 1.1. |  | If KAL O.K skip ICU load IF KAL bad re-load ICU |  |  |
| 3 | Run supplementary test and exit |  |  |  |  |

Supplementary Low Level Test (for Basic and Operational)

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 1.1 | Send invalid type (15,3) | TC(15,3) |  |  |  |
| 1.2 | Send invalid sub-type (5,15) | TC(5,15) |  |  |  |
| 1.3 | Send wrong checksum |  |  |  |  |
| 2.1 | Watchdog off | H6510 | H5345=’Off’ |  |  |
| 2.2 | RBI suspend |  | D6203 |  |  |
| 2.3 | Use RBI write to write a few words of data to 0xfc0 |  | D3920 |  |  |
| 2.4 | Read back memory and compare |  | D3921 |  |  |
| 2.5 | RBI go |  | D7204 |  |  |
| 2.6 | Watchdog on | H5510 | H5345=’On’ |  |  |

## 5.2	Operational Mode Test

Description

The function of the OM is tested in the operational mode test.

Switch on secondary rail of OM-1 and test operation.

Test operation of DPU.

Initial state

The OM (prime or redundant) is switched on and in basic mode.

Final state

The OM (prime or redundant) is switched on and in operational mode.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial state: OM switched on and in basic mode KAL power line is on |  |  |  |  |
| 1 1.1 1.2 1.3 1.4 1.5 1.6 | Switch to operational mode Load main operational code (< load_icu) Check HK packet counter has increased by 508 (prime)507 (redundant). Start Operational mode Wait 2 s. Verify operational mode started Send test packet Check TM received (verify in operational mode) | 91100 90000 H9001 90000 H1 91100 | File : ‘load ICU’ 3 Pref H5385 = ++1FC (hex) Pref H5390 = 1(hex) Pref H550 = 0,H650 = 0 Pref H5395 = ‘Safe’ |  |  |
| 2 2.1 2.2 | Verify OM-1 secondary rails on Measure OM +28V voltage and current Compare with table 2 |  | S/C procedure report: voltage and current |  |  |
| 4 4.1 4.2 4.3 4.4 4.5 4.6 4.7 4.8 4.9 4.10 | Check Housekeeping +25V secondary rail current +15V secondary rail current +11V secondary rail current +5.3V secondary rail current -5.3V secondary rail current -15V secondary rail current Heater status RBI status | 90000 | Pref H5280 = 20 mA 2 Pref H5285 = 45 mA2 Pref H5290 = 100 mA2 Pref H5295 = 480 mA2 Pref H5300= 440 mA2 Pref H5305= 55 mA2 Pref H5240 = ‘onoffoffoff’ or = ‘ononoffoff’ Pref H5340 = 1014 (hex) |  |  |
| 5 5.1 5.2 5.3 5.4 5.5 5.6 | Test DPU (start-up sequence) Send DPU reset (tc_reset_dsp) Check TC received Examine HK heart beat count Only for full functional test …. Load DPU code procedure (< load_dpu) Send Load DPUOS Check TC Received | H7201 91100 92201 90000 H7202 91100 92202 | ‘Packet accepted’ DA_DPU_BOOT_READY Pref H5410 (incrementing every 10 seconds) Ref: DPU Load Script ‘Packet received’ ‘DPU OS ready’ |  |  |
| 6 | Run supplementary test and exit |  |  |  |  |

1	Step 5.4 and 5.5 should only be followed when performing a full functional test.

2	All currents should be within the nominal range +/- 10 mA

3	There will be no telemetry command echo following this command.

| Power Rail | Voltage | Current |
| --- | --- | --- |
| +28V line | +28V +/- 2V | 1440(onoffoffoff) - 2350 (ononoffoff) |
| Keep Alive Line | +8V +/- 1V | 8mA +/- 5mA |

Table 01:	OM power consumption (with OM-1 secondaries on)

## 5.3	Filter Wheel Control Test

WARNING

The Filter wheel is a limited life time device. Each rotation of the filter wheel must be recorded in the instrument log book for the flight model.

Description

The function of the OM filter wheel operation is tested in operational mode.

Command search for coarse sensor.

Command search for datum position.

Initial state

The OM (prime or redundant) is switched on and in an operational mode.

Final state

The OM (prime or redundant) is switched on and in an operational mode with filter wheel in datum position.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial state: OM switched on and in operational mode KAL power line is on |  |  |  |  |
| 1 1.1 1.2 1.3 1.4 | Test of filter wheel coarse command Command F/W off coarse sensor (tc_fw_rel_steps 1000) (tc_fw_move) Examine HK Command F/W to coarse sensor (tc_fw_to_coarse) (tc_fw_move) Examine HK | H7606 H5600 92600 90000 H7609 H5600 92600 90000 | Pref H290 = 1000 (dec) Pref H7010 = ‘fw @ req pos’ Pref H5250 = ‘not seen’ Pref H5150 = ‘not seen’ Pref H7010 = ‘fw @ req pos’ Pref H5250 = ‘seen’ |  |  |
| 2 2.1 2.2 | Test of filter wheel datum command Command F/W to datum (tc_fw_to_datum) (tc_fw_move) Examine HK | H7608 H5600 92600 90000 | Pref H7010 = ‘fw @ req pos’ Pref H5150 = ‘seen’ Pref H5265 = 0 |  |  |
| 3 | Exit |  |  |  |  |

## 5.4	Dichroic Control Test (beam deflector)

Note

This test procedure should be reviewed by OM personnel before first execution.

The Dichroic is a limited life time device. Each movement of the dichroic wheel must be recorded in the instrument log book for the flight model instrument. A suggested movement log is given in the appendix.

Description

The function of the OM dichroic (beam deflector) is tested in engineering mode.

Command dichroic to rotate to redundant.

Command dichroic to rotate to prime.

Initial state

The OM (prime or redundant) is switched on and in an operational mode (safe, idle, science or engineering) .

Final state

The OM (prime or redundant) is switched on and in idle mode. The dichroic is at ‘prime’ interface position.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 0.1 | Initial state: OM switched on and in an operational mode KAL power line is on OM-1 secondary rails are on Ensure mode is engineering - Go to idle mode - Go to engineering mode | H9002 H9004 | H650 = 0 H650 = 0 |  |  |
| 1 1.1 1.2 1.3 1.4 | Test of dichroic command rotate to redundant Set-up dichroic to redundant (tc_dm_direction 1,0) Command dichroic to move (tc_dm_move) Examine TM Examine HK for steps in last movement | H7650 H5650 92600 90000 | Pref H370 = 1 Pref H375 = ‘max by sign’ Pref H7010 = ‘Dichroic pos’ Pref H5275 = 32 (dec) |  |  |
| 2 2.1 2.2 2.3 2.4 2.5 | Test of dichroic command rotate to prime Set-up dichroic to prime (tc_dm_direction -1,0) Command dichroic to move (tc_dm_move) Wait up to 20 seconds Examine TM (test for event) Examine HK | H7650 H5650 92600 90000 | Pref H370 = -1 Pref H375 = ‘max by sign’ Pref H7010 = ‘Dichroic pos’ Pref H5275 = -31 (dec) |  |  |
| 3 | Ensure OM is in Idle Mode | H9002 | H650 = 0 |  |  |
| 4 | Exit |  |  |  |  |

The Dichroic is a limited life time device. Each movement of the dichroic wheel must be recorded in the instrument log book for the flight model instrument.

5.5	Heater Control Test

Description

The function of the heaters is tested in an operational mode.

Each of the four main heaters are cycled on/off in turn.

Initial state

The OM (prime or redundant) is switched on and in an operational mode.

Final state

The OM (prime or rdundant) is switched on and in an operational mode.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial state: OM switched on and in operational mode KAL power line is on |  |  |  |  |
| 1 1.1 | Prepare for heater switching Disable standard control algorithm (tc_heater_stop) Examine TM Examine HK | H6670 91100 90000 | Pref H5240 = ‘offoffoffoff’ |  |  |
| 2 2.1 2.2 2.3 2.4 2.5 | Switch-on heater 1 (Main I/F) (tc_heater_config 8) Examine HK Measure S/C +28V voltage and current Compare with table 5 Switch-off heater 1 (tc_heater_config 0) Examine HK | H7660 90000 H7660 90000 | Pref H380 = ‘onoffoffoff’ Pref H5240 = ‘onoffoffoff’ S/C procedure report: voltage and current Pref H380 = ‘offoffoffoff’ Pref H5240 = ‘offoffoffoff’ |  |  |
| 3 3.1 3.2 3.3 3.4 3.5 3.6 | Switch on heater 2 (forward tube) Switch-on heater 2 (tc_heater_config 4) Examine HK Measure S/C +28V voltage and current Compare with table 5 Switch-off heater 2 (tc_heater_config 0) Examine HK | H7660 90000 H7660 90000 | Pref H380 = ‘offonoffoff’ Pref H5240 = ‘offonoffoff’ S/C procedure report: voltage and current Pref H380 = ‘offoffoffoff’ Pref H5240 = ‘offoffoffoff’ |  |  |
| 4. 4.1 4.2 4.3 4.4 4.5 4.6 | Switch on heater 3 (metering rods) Switch-on heater 3 (tc_heater_config 2) Examine HK Measure S/C +28V voltage and current Compare with table 5 Switch-off heater 3 (tc_heater_config 0) Examine HK | H7660 90000 H7660 90000 | Pref H380 = ‘offoffonoff’ Pref H5240 = ‘offoffonoff’ S/C procedure report: voltage and current Pref H380 = ‘offoffoffoff’ Pref H5240 = ‘offoffoffoff’ |  |  |
| 5 5.1 5.2 5.3 5.4 5.5 5.6 | Switch on heater 4 (secondary mirror) Switch-on heater 4 (tc_heater_config 1) Examine HK Measure S/C +28V voltage and current Compare with table 5 Switch-off heater 4 (tc_heater_config 0) Examine HK | H7660 90000 H7660 90000 | Pref H380 = ‘offoffoffon’ Pref H5240 = ‘offoffoffon’ S/C procedure report: voltage and current Pref H380 = ‘offoffoffoff’ Pref H5240 = ‘offoffoffoff’ |  |  |
| 6 6.1 | Disable heater switching Enable standard control algorithm (tc_heater_start) | H5670 |  |  |  |
| 7 | Exit |  |  |  |  |

| Heater | +28V Current |
| --- | --- |
| 1 (Main Interface) | 1440mA +/- 100mA |
| 2 (Forward tube) | 1740mA +/- 100mA |
| 3 (Metering Rods) | 1060mA +/- 100mA |
| 4 (Secondary Mirror) | 950mA +/- 100mA |
| Disable Heater Algorithm | 810mA +/- 100mA |

Table 02  Heater switching power profile

## 5.6	Image Acquisition Test - Multiple Window

WARNING

Operation of the high voltage unit must only be performed when OM personnel are present and only in a controlled environment - irreparable damage of the detector can occur.

The Filter wheel is a limited life time device. All movements of the filter wheel (such as occur in this test) must be recorded in the instrument log book for the flight model.

Description

The function of the OM Image acquisition is tested in operational mode.

A full image will be acquired and dumped to the instrument workstation. Only those areas enabled in the window table should contain events.

Initial state

The OM (prime or redundant) is switched on and in an operational mode.

OM-1 secondary rails are switched on.

Final state

The OM (prime or redundant) is switched on and in idle mode.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial state: OM switched on and in engineering mode KAL power line is on OM-1 secondary rails are on |  |  |  |  |
| 1 1.1 1.2 1.3 1.4 1.5 1.6 1.7 1.8 | Complete DPU reset Send DPU reset Check TM received Enable DPUOS code Check TM received Enable blue 1 dsp Enable blue 2 dsp Enable red dsp Init DPU Check for event | H7201 91100 92201 H7202 91100 92202 H7207 H7207 H7207 H7248 92210 | Pref H522 = ‘Blue1’ Pref H16 = ‘Enabled’ Pref H522 = ‘Blue2’ Pref H16 = ‘Enabled’ Pref H522 = ‘Red’ Pref H16 = ‘Enabled’ |  |  |
| 2 2.1 2.2 | Prepare DPU for image acquisition Ensure OM is in idle mode (tc_mode 2) Set exposure number | H9002 H7238 | Pref H650=0 Pref H530 = 1 |  |  |
| 3 3.1 3.2 3.3 3.4 3.5 3.6 3.7 | Prepare BPE for image acquisition Set acquisition mode (High Res, Windowed) Examine HK Set event threshold Specify centroid channel boundaries Load channel boundaries Wait for event ( 2 minutes) Specify window params Load window look-up table wait for event ( 1 minute ) Ensure Frame Tags disabled | H7130 90000 H7131 H7100 H5100 92100 H7110 H5110 92100 H7135 | Pref H110 = ‘Hi Res Win’ Pref H5215 = ‘Hi Res Win’ Pref H120 = 15 Table 5.6-1 H7000 = ‘cen tab load’ Table 5.6-2 H7000 = ‘win tab load’ H16 = ‘Disabled’ |  |  |
| 4 4.1 4.2 4.3 4.4 4.5 4.6 | Place Filter Wheel in Blocked Position. Ensure coarse and fine sensors have correct values (tc_fw_coarse_sensor_current 4 tc_fw_fine_sensor_current 9) Specify filter wheel to datum Move it there. Await Event Specify move to blocked (Filt. = 0) Move it there Await Event | H7600 H7601 H7608 H5600 92600 H7604 H5600 92600 | H300 = 4 (56 mA) H310 = 9 (25 mA) H7010 = ‘fw @ req pos’ H280 = ‘Blocked’ H7010 = ‘fw @ req pos’ |  |  |
| 5 5.1 5.2 5.3 5.4 5.5 5.6 5.7 5.8 5.9 5.10 5.11 5.12 | If using Redundant Chain, go to next section. Set up High Voltages (Prime) (Warning, these are hazardous commands) Specify Vmcp23 Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp . Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. | H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 | H150=’mcp23’,H160=1400, H175=’off’ H7000 = ‘mcp23 ok’ H150=’mcp1’,H160=550, H175=’off’ H7000 = ‘mcp1 ok’ H150=’cathode’,H160=2, H175=’off’ H7000 = ‘cathode ok’ H150=’cathode’,H160=396, H175=’off’ H7000 = ‘cathode ok’ H150=’mcp1’,H160=520, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’,H160=1700, H175=’off’ H7000 = ‘mcp23 ok’ |  |  |
| 6 6.1 6.2 6.3 6.4 6.5 6.6 6.7 6.8 6.9 6.10 6.11 6.12 | If using Prime chain, go to next section Set up High Voltages (Redundant) (Warning, these are hazardous commands) Specify Vmcp23 Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp . Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. | H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 | H150=’mcp23’,H160=1500, H175=’off’ H7000 = ‘mcp23 ok’ H150=’mcp1’,H160=500, H175=’off’ H7000 = ‘mcp1 ok’ H150=’cathode’,H160=2, H175=’off’ H7000 = ‘cathode ok’ H150=’cathode’,H160=390, H175=’off’ H7000 = ‘cathode ok’ H150=’mcp1’,H160=550, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’,H160=1850, H175=’off’ H7000 = ‘mcp23 ok’ |  |  |
| 7 7.1 7.2 7.3 7.4 7.5 7.6 7.7 | Acquire image. Ensure we are in engineering mode (tc_mode 4). Turn on Flood LED (level =2). Set DPU frame time to 5 mins. (ic_set_frame_time) Start sending events to DPU. Start full image eng integration (tc_enbl_eng 1,4). Wait 5 minutes to acquire data Await Event (DA_COMPLETE_EXP). Stop sending events to DPU (failsafe command). Disable Flood LED (level=0) | H9004 H7134 H7236 H5130 H7254 92209 H6130 H7134 | H650 = 0 H1040 = 2 (4 uA) H710 = 300000 H880 = ‘Full hi res’ H1040 = 0 (0 mA) |  |  |
| 8 8.1 8.2 8.3 8.4 8.5 8.6 | Ramp down HV (Warning, these commands are hazardous) Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. | H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 | H150=’cathode’,H160=0, H175=’off’ H7000 = ‘cathode ok’ H150=’mcp1’, H160=0, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’,H160=0, H175=’off’ H7000 = ‘mcp23 ok’ |  |  |
| 9 | Await Dumping of Engineering Image (2.5 hours) |  |  |  |  |
| 10 | Exit |  |  |  |  |

| Description | Pref | Value |
| --- | --- | --- |
| Enable/Disable Verification | H10 | ‘Enable’ |
| X Centroid Table Boundary 0 | H20 | -1000 |
| X Centroid Table Boundary 1 | H21 | -750 |
| X Centroid Table Boundary 2 | H22 | -500 |
| X Centroid Table Boundary 3 | H23 | -250 |
| X Centroid Table Boundary 4 | H24 | 0 |
| X Centroid Table Boundary 5 | H25 | 250 |
| X Centroid Table Boundary 6 | H26 | 500 |
| X Centroid Table Boundary 7 | H27 | 750 |
| X Centroid Table Boundary 8 | H28 | 1000 |
| Y Centroid Table Boundary 0 | H30 | -1000 |
| Y Centroid Table Boundary 1 | H31 | -750 |
| Y Centroid Table Boundary 2 | H32 | -500 |
| Y Centroid Table Boundary 3 | H33 | -250 |
| Y Centroid Table Boundary 4 | H34 | 0 |
| Y Centroid Table Boundary 5 | H35 | 250 |
| Y Centroid Table Boundary 6 | H36 | 500 |
| Y Centroid Table Boundary 7 | H37 | 750 |
| Y Centroid Table Boundary 8 | H38 | 1000 |

Table 5.6-1 Centroid Table Boundaries

| Description |  | Pref | Value |
| --- | --- | --- | --- |
| Enable/Disable Verification = 1/0 |  | H10 | ‘Enable’ |
| No of Windows |  | H40 | 2 |
| Xlow | Parameters for | H50 | 64 |
| Ylow | Window 1 | H51 | 15 |
| Xsize | (Units are CCD pixels) | H52 | 64 |
| Ysize |  | H53 | 64 |
| Xlow | Parameters for | H54 | 216 |
| Ylow | Window 2 | H55 | 77 |
| Xsize | (Units are CCD pixels) | H56 | 64 |
| Ysize |  | H57 | 64 |
|  | All remaining parameters | H58-H109 | 0 |

Table 5.6-2 Window Parameters

## 5.7	Image Acquisition Test - Full Frame

WARNING

Operation of the high voltage unit must only be performed when OM personnel are present and only in a controlled environment - irreparable damage of the detector can occur.

The Filter wheel is a limited life time device. All movements of the filter wheel (such as occur in this test) must be recorded in the instrument log book for the flight model.

Description

The function of the OM Image acquisition is tested in engineering mode.

A full frame image will be acquired and dumped to the instrument work station.

Initial state

The OM (prime or redundant) is switched on and in an operational mode (safe, idle, science or engineering).

OM-1 secondary rails are switched on.

Final state

The OM (prime or redundant) is switched on and in engineering mode.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial state: OM switched on and in an operational mode KAL power line is on OM-1 secondary rails are on |  |  |  |  |
| 1 1.0 1.1 1.2 1.3 1.4 1.5 1.6 1.7 | Complete DPU reset Send DPU reset Check TM received Enable DPUOS code Check TM received Enable blue 1 dsp Enable blue 2 dsp Enable red dsp Init DPU Check for event | H7201 91100 92201 H7202 91100 92202 H7207 H7207 H7207 H7248 92210 | Pref H522 = ‘Blue1’ Pref H16 = ‘Enabled’ Pref H522 = ‘Blue2’ Pref H16 = ‘Enabled’ Pref H522 = ‘Red’ Pref H16 = ‘Enabled’ |  |  |
| 2 2.1 2.2 | Prepare DPU for image acquisition Ensure OM is in idle mode (tc_mode 2) Set exposure number | H9002 H7238 | Pref H650=0 Pref H530 = 2 |  |  |
| 3 3.1 3.2 3.3 3.4 3.5 3.6 3.7 | Prepare BPE for image acquisition Set acquisition mode (High Res, Full Frame) Examine HK Set event threshold Specify centroid channel boundaries Load channel boundaries Wait for event ( 2 minutes) Specify window params Load window look-up table wait for event ( 1 minute ) Ensure Frame Tags disabled | H7130 90000 H7131 H7100 H5100 92100 H7110 H5110 H7135 | Pref H110 = ‘Hi Res Full’ Pref H5215 = ‘Hi Res Full’ Pref H120 = 15 Table 5.6-1 H7000 = ‘cen tab load’ Table 5.7-1 H7000 = ‘win tab load’ H16 = ‘disabled’ |  |  |
| 4 4.1 4.2 4.3 4.4 4.5 4.6 | Place Filter Wheel in Blocked Position. Ensure coarse and fine sensors have correct values (tc_fw_coarse_sensor_current 4 tc_fw_fine_sensor_current 9) Specify filter wheel to datum Move it there. Await Event Specify move to blocked (Filt. = 0) Move it there Await Event | H7600 H7601 H7608 H5600 92600 H7604 H5600 92600 | H300 = 4 (56 mA) H310 = 9 (25mA) H7010 = ‘fw @ req pos’ H280 = ‘Blocked’ H7010 = ‘fw @ req pos’ |  |  |
| 5 5.1 5.2 5.3 5.4 5.5 5.6 5.7 5.8 5.9 5.10 5.11 5.12 | If using Redundant Chain, go to next section. Set up High Voltages (Prime) (Warning, these are hazardous commands) Specify Vmcp23 Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp . Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. | H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 | H150=’mcp23’,H160=1400, H175=’off’ H7000 = ‘mcp23 ok’ H150=’mcp1’,H160=550, H175=’off’ H7000 = ‘mcp1 ok’ H150=’cathode’,H160=2, H175=’off’ H7000 = ‘cathode ok’ H150=’cathode’,H160=396, H175=’off’ H7000 = ‘cathode pk’ H150=’mcp1’,H160=520, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’,H160=1700, H175=’off’ H7000 = ‘mcp23 ok’ |  |  |
| 6 6.1 6.2 6.3 6.4 6.5 6.6 6.7 6.8 6.9 6.10 6.11 6.12 | If using Prime chain, go to next section Set up High Voltages (Redundant) (Warning, these are hazardous commands) Specify Vmcp23 Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vcathode Ramp . Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. | H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 | H150=’mcp23’,H160=1500, H175=’off’ H7000 = ‘mcp23 ok’ H150=’mcp1’,H160=500, H175=’off’ H7000 = ‘mcp1 ok’ H150=’cathode’,H160=2, H175=’off’ H7000 = ‘cathode ok’ H150=’cathode’,H160=390, H175=’off’ H7000 = ‘cathode’ H150=’mcp1’,H160=550, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’,H160=1850, H175=’off’ H7000 = ‘mcp23 ok’ |  |  |

| 7 7.1 7.2 7.3 7.4 7.5 7.6 7.7 | Acquire image. Ensure we are in engineering mode (tc_mode 4). Turn on Flood LED (level =2). Set DPU frame time to 5 mins. (ic_set_frame_time) Start sending events to DPU. Start full image eng integration (tc_enbl_eng 1,4). Wait 5 minutes to acquire data Await Event (DA_COMPLETE_EXP). Stop sending events to DPU (failsafe command). Disable Flood LED (level=0) | H9004 H7134 H7236 H5130 H7254 92209 H6130 H7134 | H650 = 0 H1040 = 2 (4 uA) H720 = 300000 H880 = ‘Full hi res’ H1040 = 0 uA |  |  |
| --- | --- | --- | --- | --- | --- |
| 8 8.1 8.2 8.3 8.4 8.5 8.6 | Ramp down HV (Warning, these commands are hazardous) Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. | H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 | H150=’cathode’,H160=0, H175=’off’ H7000 = ‘cathode ok’ H150=’mcp1’,H160=0, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’, H160=0, H175=’off’ H7000 = ‘mcp23 ok’ |  |  |
| 9 | Await Dumping of Engineering Image (2.5 hours) |  |  |  |  |
| 10 | Exit |  |  |  |  |

| Description |  | Pref | Value |  |
| --- | --- | --- | --- | --- |
|  |  |  | Prime | Redundant |
| Enable/Disable Verification = 1/0 |  | H10 | ‘Enable’ | ‘Enable’ |
| No of Windows |  | H40 | 1 | 1 |
| Xlow | Parameters for | H50 | 52 | 64 |
| Ylow | Window 1 | H51 | 27 | 15 |
| Xsize | (Units are CCD pixels) | H52 | 256 | 256 |
| Ysize |  | H53 | 256 | 256 |
|  | All remaining parameters | H54-H109 | 0 | 0 |

Table 5.7-1 Window Parameters

## 5.8	Switch-off

Description

Mode and state are checked and the appropriate procedure is followed to switch-off in an orderly way.

Initial state

The OM (prime or redundant) is switched on.

Final state

The OM (prime or redundant) is switched off. The KAL power lines may be left on.

| Step | Action | TC/TM | Remarks | Pass | Notes |
| --- | --- | --- | --- | --- | --- |
| 0 | Initial state: OM switched on KAL power line is on |  |  |  |  |
| 1 | Check OM mode if mode = Operational = Basic |  | goto step 2 goto step 3 |  |  |
| 2. 2.1 2.2 2.3 2.4 2.5 2.6 | If HV enabled: Ramp down HV (Warning, these commands are hazardous) Specify Vcathode Ramp. Perform Ramp. Await Event. Specify Vmcp1 Ramp. Perform Ramp. Await Event. Specify Vmcp23 Ramp. Perform Ramp. Await Event. Examine HK | 90000 H7140 H5140 92100 H7140 H5140 92100 H7140 H5140 92100 90000 | H5145 = 1 = ‘Enabled’ H150=’cathode’,H160=0, H175=’off’ H7000 = ‘cathode ok’ H150=’mcp1’,H160=0, H175=’off’ H7000 = ‘mcp1 ok’ H150=’mcp23’,H160=0, H175=’off’ H7000 = ‘mcp23 ok’ H5145=’Disabled’ |  |  |
| 3.1 3.2 | Switch off OM-1 secondary rails Disable heater algorithm | H6690 H6670 |  |  |  |
| 4 | Switch-off +28V power |  | S/C procedure report: voltage and current |  |  |
| 5 | Switch-off KAL |  | S/C procedure report: voltage and current |  |  |
| 6 | Exit |  |  |  |  |

Appendix

Parameter reference Table

| Pref ID | HK Monitor |
| --- | --- |
| H5150 | Fine Pick-Off Sensor |
| H5215 | Acquisition Mode |
| H5240 | Commanded Heater Status |
| H5250 | Coarse Sensor Status |
| H5265 | Filter Wheel Position Counter |
| H5275 | S/w Dichroic Position Counter |
| H5280 | Switched +25V Current |
| H5285 | Switched +15V Current |
| H5290 | Switched +11V Current |
| H5295 | Switched +5.3V Current |
| H5300 | Switched -5.3V Current |
| H5305 | Switched -15V Current |
| H5320 | ICB Status |
| H5340 | RBI Status |
| H5345 | RBI Config |
| H5385 | Good TC |
| H5390 | Bad TC |
| H5395 | OM State |
| H5410 | Heartbeat Count |

Table 03:	Pref look-up Table

Dichroic Rotation Log

| - ÿ | - ÿ |
| --- | --- |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |
| - ÿ | - ÿ |